TECHNICAL NOTE xx fibroscan 402 installation xl patch

Written by Nicolas SOLLIER

# In which case you need to use this procedure

When you need to install XL patch on Fibroscan FS402

# Required equipment

-USB Stick empty

-XL PATCH 402.zip

# Instruction

-Unzip XL patch 402.zip and copy paste all files on memory stick

-Turn ON Fibroscan 402

-Then Go to the configuration menu and select log in as Biomedical Engineer

Password: maintenance   (all lower case)

-Press configuration Menu

-Select Biomedical Engineer

Password:  maintenance and then press OK

You will get the screen below and then press service tab

-Press service tab

***NOW PLUG IN USB STICK XL PATCH 2

Press the “Launch Program” and follow the following screen instructions

-Select FS402_Patch_1.0.0.0_signed.exe and then press OK

-Press OK

-Press OK, wait and reboot the system