## Slide 1
fixed image
Reinstall the software
( see page 34 of E200M006_5 manual)
Try another motherboard
 if needed reconnect or replace the mother board (see page 22 of E200M006_5 technical manual)
Connect an external mouse to see if the touchpad reponds well
 if needed replace the screen (see page 32 of E200M006_5 technical manual)
unsolved
unsolved
solved
ACTIONS
Request a new compact flash to ECHOSENS service (service@echosens.com) and replace it in the device ( see page 23 of E200M006_5 manual)