## Slide 1
Symptom : No printout
Check voltages on the power cable (see page 26 of E200M006_5 technical manual)
 If needed change the power board ( see (page 28 of E200M006_5 manual)
Check that power and COM cables are connected  If needed reconnect or replace those and turn ON (if disconnected reglue as stated in critical updates)
Check that paper is in the right position
 If needed change the paper ( see page 16 of E200M006_5 technical manual)
 Replace the printer see page 17 of E200M006_5 technical manual
unsolved
unsolved
unsolved
solved
ACTIONS