## Slide 1
Symptom : the instrument does not switch on
Check the wall outlet value
 If needed use another outlet
Check the Mains filter fuse
Check the Mains filter outlet voltage value
 If needed replace it
Check the motherboard power supplies (page 27 of E200M006_5 technical manual)
Check the mother board power supplies cable
 If needed replace the power supply (page 29 of E200M006_5 technical manual) board or reconnect the cable
Check ON/OFF button connection to mother board  if needed reconnect or replace the button
 Replace the motherboard (see page 22 of E200M006_5 technical manual
unsolved
unsolved
unsolved
unsolved
solved
ACTIONS