## Slide 1
the instruments turn on but nothing on the screen
Connect a screen to the rear panel DVI connector
 If it comes on check the screen connector on the motherboard (page 22 of E200M006_5 technical manual
Check the inverter board
 If needed replace it (page 33 of E200M006_5technical manual
Try with another screen
If needed replace the screen (page 33 of E200M006_5technical manual
→Replace the motherboard (see page 22 of E200M006_5 technical manual)
unsolved
unsolved
unsolved
solved
ACTIONS