## Slide 1
probe not connected error message
Check that the probe is well connected
 Change the probe or contact ECHOSENS service
Check that probe connector pins are not damaged
 Rebend the pins or send the probe to ECHOSENS service
Check the voltage value of the REGANA board (step 2 to 4)
 Change the elastography module (see page 42 of E100M011manual or movie 3)
unsolved
unsolved
solved
ACTIONS
Reboot the equipment in case the device is in a temporary unstable state