## Slide 1
Symptom:the Fibroscan does not switch off from the start button
Check the fuses on the main switch
 If needed change the main switch
Check the start button connector and the ATX remote connector
 If needed replace the ON/OFF button
Check the mother board voltage value (see page 31 of of E100M011
Manual)
 If needed replace mother board (see page 32 of E100M011 technical manual or movie 6)
Check the power supply cable
Reconnect it or change power module (see page 48 of E100M011 manual)
unsolved
unsolved
unsolved
solved
ACTIONS