## Slide 1
Symptom : the instrument does not switch on
Check the wall outlet value
 If needed use another outlet
Check the Mains filter fuse
Check the Mains filter outlet voltage value
 If needed replace it
Check the motherboard power supplies (page 50 of E100M011 manual)
Check the mother board power supplies cable
 If needed replace the power supply (see (page 48 of E100M011 manual)
or reconnect the cable
Check ON/OFF button connection to mother board or remote control board
 if needed reconnect those
unsolved
unsolved
unsolved
solved
ACTIONS