## Slide 1
the instruments turn on but nothing on the screen
Connect a screen to the front panel VGA connector
 If it comes on check the screen connector on the motherboard (see page 29 of E100M011manual
)
Check the inverter board
 If needed replace it (see page 34 of E100M011 manual or movie 8)
Try with another screen
if needed reconnect it or replace it (see page 34 of E100M011 manual or movie 8)
→ if needed reconnect or replace the mother board ( see page 32 of E100M011 manual or movie 6)
unsolved
unsolved
unsolved
solved
ACTIONS