## Slide 1
Symptom : No printout
Check that power cable is connected and that the printer is ON
 If needed reconnect and turn ON
Check the paper tray and printer cartridge
 If needed refill with paper and install new cartridge (see printer manual)
Check that the printer is listed in the preinstalled ones
Check that the printer works connected to a computer
If it is a network printer try to PING it
 Delete the printer from FS and reinstall it
 Replace the printer or try another model
unsolved
unsolved
unsolved
solved
ACTIONS
unsolved
 Change PV3 (page 23 of E320M003-2 manual)
 Replace the printer or try another model
unsolved