## Slide 1
theSymptom : no shearwave generated on low battery
If battery charge is inferior to 5% the device will be on but the probe will not generate any shearwave  install patch H2,0,6, the FS530 will not allow any exam on low battery
Note that FS530 running OS G3,1,22 will have a threshold of 20ù no need for modification
solved
ACTIONS