|  |  |
| --- | --- |
|  | Checking-List 430 |

|  |  |
| --- | --- |
| REF. DOCUMENT : | QF217 |
|  | Nom | Fonction | Date | Signature |
| AUTEUR | Nicolas SOLLIER | Technicien SAV |  |  |
| VERIFICATEUR | Nathalie LUONG | Chargée Qualité |  |  |
| APPROBATEUR | Olivier BOSCHAT | Responsable SAV |  |  |
|  |
| DIFFUSION : | ECHOSENS, DISTRIBUTEURS |
|  |
| MODIFICATIONS : |
| Date | Indice | Nature de la modification | Par |
| 14/06/2017 | 1 | Création du document | OB |
| 16/08/2017 | 2 | Suppression de 2 tests électriques :- Protective earth impedance - Breakdown on insulation | OBO |
| 27/09/2017 | 3 | Case Network Output : ajout de la case N/ARemise en forme des informations produitSuppression des tests liés au Wifi/Bluetooth (suite au retrait du module correspondant)Ajout du schéma dans la partie « Power supply »Ajout des parties « Probe 2 » et « Probe 3 »Ajout d’une cartouche de signature à la fin du rapport | NSO |

|  |  |  |
| --- | --- | --- |
| Form REF. | QF217.3\_SOP QP052 | 2 pages |

|  |
| --- |
| **Product Informations** |
| **S/N** |  | **Man. /Dist.** |  |
| **Probe S/N** |  |  |  |
| **Customer location** |  | **Customer name** |  |

|  |  |
| --- | --- |
| **Service location** | [ ]  Echosens [ ]  On site ………………………………… [ ]  Other:……………………………..  |
| **Service date** |   | **Service reference** | SR/ |

|  |
| --- |
| **Technical Note Applicable** |
| 52-Outer Screws Exchange | [ ]  OK [ ]  NOT OK  |

|  |
| --- |
| **Software** |
| Update software | [ ]  OK [ ]  NOT OK [ ]  N/A | Software version = |
| Check Log files | [ ]  OK [ ]  NOT OK |
| Check import/export data | [ ]  OK [ ]  NOT OK |
| Magic fix performed with success (CHKDK) | [ ]  OK [ ]  NOT OK |

|  |
| --- |
| **Used instruments** |
| **Manufacturer** | **Model** |  **Description** | **MET (Echosens only)****Or S/N (for distributors)** | **Calibration due date** |
|  |  |  |  |  |
|  |  |  |  |  |



|  |
| --- |
| **Power Supply**  |
| Output + 19 Vdc ±0,5Vdc (Pin 1-2)= V | [ ]  OK [ ]  NOT OK |
| Battery voltage (Between + and -)= V  | [ ]  OK [ ]  NOT OK |

|  |
| --- |
| **Inner Electronic** |
| Check every connector | [ ]  OK [ ]  NOT OK |
| Front LED indication  | [ ]  OK [ ]  NOT OK |
| PV3 LED status | [ ]  OK [ ]  NOT OK |

|  |
| --- |
| **Outputs** |
| Test on/off button (led on) | [ ]  OK [ ]  NOT OK |
| Check the Touchpad | [ ]  OK [ ]  NOT OK |
| Test USB ports | [ ]  OK [ ]  NOT OK Comment: |
| Test probe connector 1 | [ ]  OK [ ]  NOT OK |
| Test probe connector 2 | [ ]  OK [ ]  NOT OK |
| Test network output | [ ]  OK [ ]  NOT OK [ ]  N/A |

|  |
| --- |
| **Probe 1** |
| Blue LEDS of the probe | [ ]  OK [ ]  NOT OK |
| Check the probe memory | Date of next calibration :………............. | [ ]  OK [ ]  NOT OK |
| Acquisition Test  | [ ]  OK [ ]  NOT OK |
| Check the information contact on calibration message  | [ ]  OK [ ]  NOT OK |

|  |
| --- |
| **Probe 2** **[ ]  N/A** |
| Blue LEDS of the probe | [ ]  OK [ ]  NOT OK |
| Check the probe memory | Date of next calibration :………............. | [ ]  OK [ ]  NOT OK |
| Acquisition Test  | [ ]  OK [ ]  NOT OK |
| Check the information contact on calibration message  | [ ]  OK [ ]  NOT OK |

|  |
| --- |
| **Probe 3** **[ ]  N/A** |
| Blue LEDS of the probe | [ ]  OK [ ]  NOT OK |
| Check the probe memory | Date of next calibration :………............. | [ ]  OK [ ]  NOT OK |
| Acquisition Test  | [ ]  OK [ ]  NOT OK |
| Check the information contact on calibration message  | [ ]  OK [ ]  NOT OK |

|  |
| --- |
| **Class 1 Type B safety test** |
| Date………………Tested By……………….…..…… | Earth leakage current (Equipment leakage)<500µA [ ]  Pass    [ ]  Fail |
|  |
| IEC 62353:2008 |

|  |
| --- |
| **Frame** |
| State of device (cf. picture below) | [ ]  OK [ ]  NOT OK |
| Labelling | [ ]  OK [ ]  NOT OK |

**FRAME CONTROL**

****

|  |  |
| --- | --- |
| Operator Code: | Signature: |