## Slide 1
Technical training

## Slide 2
Focal One softwares

## Slide 3
FOnePowerOff:
Initialization and power off trought MEP board.
FOneTherapy:
Treatment software.
FOneSetup:
Setup and maintenance software.
FOneXview:
Treatment images and software.
USScanner_TcpIp:
Ultrasound imaging and communication settings.
Softwares installed on computer

## Slide 4
MEP_Fone:
MEP/Extension and accessory board software.
GF_One:
Cooling system software.
MOT_Fone:
Motorization software.
Probe_Fone:
Probe software.
Ampli_Fone:
Amplifier software.
Softwares not installed on computer

## Slide 5: Softwares Overview
Focal One computer
(Windows 7)
FOneTherapy
FOneSetup
PowerOff
Matrox software
Mains electronics boards
MEP_FOne
Amplifier
Ampli_FOne
Focal One computer
(Acquisition board)
USScanner_TcpIp
FOneXView
Ultrasound
DVI / VGA
Ethernet
USB
USB
Cooling unit
GF_One
Probe’s motors
MOT_FOne
DF Probe
Probe_FOne
CAN Bus
RS232
CAN Bus

## Slide 6
FOneTherapy

## Slide 7
Treatment software.
Used by hospital staff.
Boots automatically when windows starts on ‘Focal One’ user.
No adjustments with this software. It used technical paramaters calibrated with FOneSetup.
Hospital password needed to access to hospital data base.
1 hospital = 1 data base.
2 levels available :
  Level 0 = Normal level of use.
  Level 1 (Mabla123) = Only available for ‘Edap’ hospital.
FOneTherapy

## Slide 8: FOneTherapy – Treatment review

## Slide 9
FOneTherapy – Treatment screen

## Slide 10
FOneSetup

## Slide 11
Maintenance software.
Doesn’t start automatically.
Trained technician use only .
No treatment with this software, only service function:
  Calibrations.
  Settings.
  Parameters reading.
  2 levels available
  Level 1 (Mabla123) = Normal level of use.
  Level 2 = Only available for software team.
FOneSetup

## Slide 12
Password
Init
Main page
Technical file
Maintenance
Language
Level
QUIT
FOneSetup

## Slide 13
Maintenance
MEP Digital I/O
MEP Analog I/O
Probe holder
Cooling
Firing and power
Video
Versions
FOneSetup – Maintenance Menu

## Slide 14: FOneSetup – Technical file Menu
Technical file
MEP Config
Motors
Cooling
Probe
Amplifier
Protocols
Imaging
PRINT
Options config
EEPROM
EEPROM
EEPROM
EEPROM

## Slide 15
FOneXview

## Slide 16
Imaging software.
Display MRI imaging and ultrasound imaging.
Fusion is possible with this software.
Display old treatments (ex Ablaview).
Can display biopsie reports.
FOneXView

## Slide 17
FOneXView

## Slide 18
FOnePowerOff
USScanner_TCPIP

## Slide 19: FOnePowerOff
This part is displayed only in administator session.

## Slide 20: USScanner_TcpIp / Echo+