## Slide 1: Technical training
Maintenance

## Slide 2: Errors

## Slide 3: Definition of alarm in Focal One
For all software modules, the triggered alarms are coded as follows (only numbers:
W
.
.
X
YY
ZZ
Exemple: 2.5.1102 => Probe temperature is too high : XX.X

## Slide 4: Definition of alarm in Focal One
W
.
.
X
YY
ZZ
Were the number mean:
Level of alarm severity, 3 levels are possible:
  Level 1: This level corresponds to a user information that is displayed without an explicit number, and with a level 1 alarm icon. It indicates an event that is easy to solve. The alarm is still logged in the error tracking file.
  Level 2: Non-blocking alarm: Possible actions are "Retry, Continue or Ignore, Cancel or Restart". It is displayed with the error number and a level 2 alarm icon.
  Level 3: Blocking alarm: No possibility to continue treatment, the only option is abandonment. It is displayed with the error number and a level 3 alarm icon.
W

## Slide 5: Definition of alarm in Focal One
W
.
.
X
YY
ZZ
Were the number mean:
Software progress phase at the time of the alarm:
  0: Login phase
  1: Review phase of the treatment database
  2: Initialization phase of the treatment
  3: Phase selection and editing patient data and treatment
  4: Insertion phase of the probe in the patient
  5: Volume acquisition phase
  6: Definition phase of the area to be treated (treatment limits)
  7: Definition phase of lesions on cross sections
  8: Treatment phase (shooting generation)
  9: Zone End Phase (data backup, and move)
  10: End of treatment phase
  100: special phase that groups all the previous phases, but the error is generated with the use of the software FoneSetup
X

## Slide 6: Definition of alarm in Focal One
W
.
.
X
YY
ZZ
Were the number mean:
Alarm group: the group refers to a specific module:
  10: Alarms related to the power generator
  11: Cold group alarms
  12: General Security Alarms
  13: Alarms related to the detection of the rectal wall
  14: Alarms related to the sensor motor assembly (MOT)
  15: Alarms related to the sensor module (data control)
  16: Alarms related to the ultrasound system
  17: Acquisition Card Alarms
  18: Alarms related to file errors
  19: Alarms related to the management of the Focalpak code
  50: Alarms related to general functions of the MEP_Fone
  51: Reasons for canceling treatment
  52: Alarms related to hard disk management
YY

## Slide 7: Definition of alarm in Focal One
W
.
.
X
YY
ZZ
Were the number mean:
Number of the alarm in the group.
ZZ
Alarm | Error message
1001 | Power too high, higher than 10W
1002 | Power too high, higher than 10W
1003 | Active power < 80% of the required power
1004 | Active power > 120% of the required power
1005 | Reflected power higher than 20% the live power
1006 | At least four of the 16 channels are in error (difference> 20%)
1007 | Temperature alarm of the generator power
1008 | Phase of signal out of range
1009 | The initialization of the communication with the generator failed
1010 | Not receiving the stop command path
1011 | Impossible to communicate with amplifier
1012 | Error status integrity amplifier
1013 | Internal error in the amplifier
1101 | The probe temperature is too low
1102 | The probe temperature is too high
1103 | Cooling system too low < 2°C
1104 | Cooling system too high > 10°C
1105 | No communication with cooling system board
1106 | Integrity failure on cooling system
1107 | Wrong status with cooling system
1108 | The cooling system cannot communicate with the probe
1201 | 3D arm position is different from the stored position (± 1mm)
1202 | Patient movement detection
1250 | No 24V: Check if the emergency stop is released
1301 | Rectal wall is not detected
1302 | Difference between the recorded rectal wall distance -measured distance during treatment > max deviation before correction
1303 | Difference between the recorded rectal wall distance -measured distance during treatment > max deviation after adjustment

## Slide 8: Motorization errors message
X axis errors | 
1410 | Communication timeout with motors (error number associated with the motor number)
1411 | Requested X motor position is out of range
1412 | X Motor displacement timeout
1413 | Requested encoder position out of range
1414 | Difference X encoder position / control position > threshold for the current axis.
1415 | " X motor still running"
1416 | X encoder position > to the target position while moving
1417 | Centering error for X motor
1418 | Number of pulses error for X motor
1419 | No movement is detected on X motor
T axis erros | 
1430 | Communication timeout with a motor (error number associated with the motor number)
1431 | Requested T motor position is out of range
1432 | T Motor displacement timeout
1433 | Requested encoder position out of range
1434 | Difference T encoder position / control position > threshold for the current axis.
1435 | " T motor still running"
1436 | T encoder position > to the target position while moving
1437 | Centering error for T motor
1438 | Number of pulses error for T motor
1439 | No movement is detected on T motor
Y axis errors | 
1420 | Communication timeout with a motor (error number associated with the motor number)
1421 | Requested Y motor position is out of range
1422 | Y Motor displacement timeout
1423 | Requested encoder position out of range
1424 | Difference X encoder position / control position > threshold for the current axis.
1425 | " Y motor still running"
1426 | Y encoder position > to the target position while moving
1427 | Centering error for Y motor
1428 | Number of pulses error for Y motor
1429 | No movement is detected on Y motor
Z axis erros | 
1440 | Communication timeout with engine (error number associated with the engine number)
1441 | Requested Z motor position is out of range
1442 | Z Motor displacement timeout
1443 | Requested encoder position out of range
1444 | Difference Z encoder position / control position > threshold for the current axis.
1445 | " Z motor still running"
1446 | Z encoder position > to the target position while moving
1447 | Centering error for Z motor
1448 | Number of pulses error for Z motor
1449 | No movement is detected on Z motor
General errors | 
1401 | No communication with stepper motor boards
1402 | Integrity error in a stepper motor board
1403 | Fault status (authorized or not) on a motor

## Slide 9: Motorization errors message
1501 | Integrity error in probe data board (CRC) or bad values
1502 | No communication with the probe board
1503 | Probe is not calibrate with this generator
1504 | Half probe is not calibrate with this generator
1505 | Unable to move the probe configuration into full probe
1506 | Unable to move the probe configuration into half-probe
1508 | The power calibration file in probe with the generator of the machine is not present or not readable
1601 | Communication problem with the ultrasound
1602 | Problem when launching the ultrasound software
1603 | Wrong CRC ultrasound software
1604 | Unable to move the probe in run mode
1605 | Unable to retrieve the scale ultrasound
1606 | Unable to move the ultrasound image in the right direction (no flip)
1607 | The ultrasound probe is disconnected
1608 | The ultrasound is off
1701 | Problem with the initialization of ultrasound video signal
1702 | Problem with the acquisition of ultrasound video signal
1801 | Unable to backup information
1802 | Unable to open file
1803 | Writing error in data patient file
1804 | Reading error of the data patient file and impossible to recover the backup file
1805 | Integrity error of a critical file
1806 | Integrity Failure in FocalPak file
1807 | Writing error in FocalPak file
1808 | Unable to open the data sheet
1809 | Problem when saving the data sheet
1810 | Problem when saving the data sheet backup
1811 | Unable to create hospital directory
1812 | Unable to open file FocalPak
1901 | FocalPak incorrect code
1902 | FocalPak code already used
1903 | FocalPak Duration of use exceeded
1904 | Invalid code for the selected hospital
5001 | No communication with the MEP board
5002 | Integrity error on the MEP board (CRC program, CRC EEPROM, Watchdog, Power, Processor, Watchdog 2)
5003 | No 24V on the MEP board
5004 | Unable to open the communication port with the CRM
5101 | Entering patient information interrupted by user
5102 | Longitudinal volume acquisition interrupted by user
5103 | Definition of the treatment area limit is interrupted by user
5104 | Transversal localization stopped by user
5105 | treatment interrupted by user
5106 | Failed to initialize
5107 | Treatment interrupted for some reason

## Slide 10: Definition of alarm in Focal One
All alarms are sent back to the therapy software or FOneSetup which saves them in a common text file: “Error.txt” :
HOSPITAL N°PATIENT DATE J/M/A TIME H:M:S NUM ERROR MESSAGE
--------- ---------- ---------- --------- -------
EDAP N° ---- 16/03/17 15:38:55 2.0.5004 Impossible to open COM Port for MEP
N° ---- 16/03/17 16:49:26 2.0.1904 Accès refusé ! Mot de passe invalide pour cet hôpital
N° ---- 06/04/17 09:32:31 2.0.1904 Access denied! Code invalid for this hospital
EDAP N° 0008 06/04/17 09:35:52 2.5.1102 Probe temperature is too high : 22.5
EDAP N° 0008 06/04/17 09:39:33 1.5.5102 Treatment cancelled at acquisition phase
EDAP N° 0009 06/04/17 09:47:05 2.5.1102 Probe temperature is too high : 22.8
EDAP N° 0009 06/04/17 10:01:21 1.9.5104 Treatment cancelled at transversal localization phase
EDAP N° ---- 06/04/17 15:20:11 3.2.1503 The probe has not been calibrated on this machine
EDAP N° ---- 06/04/17 15:23:26 2.2.1107 Status error on Cooling System: Second temperature sensor disconnected
EDAP N° ---- 06/04/17 15:49:57 2.2.1107 Status error on Cooling System: Second temperature sensor disconnected
EDAP N° 0010 06/04/17 16:24:30 1.8.1302 Rectal wall position error Measured rectum wall position (30.2 mm) must be in the range 21.8 mm to 29.8 mm
EDAP N° ---- 07/04/17 08:55:48 2.3.1901 Incorrect Focalpak code