## Slide 1: Technical Training: Lesson 8 Maintenance

## Slide 2
TREATMENT STEP BY STEP
LESSON GOALS:
Know preventive maintenance documents
Know Electrical Safety Test documents
Know Focal One error management
Know Focal One error code
Know how to perform software update.

## Slide 3: Preventive maintenance
Preventive maintenance must be perform every 6 month.
Expected preventive maintenance duration as follow:
To perform preventive maintenance, follow instruction of latest version of the instruction
TMS511991 or INS-367.
Report result in checklist TMS510642 or FMU-431
 | 1 Probe | 2 Probes
Beginner | 1.5 days | 2 days
Confirmed | 1 day (longer than usual) | 1.5 days
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 4: Electrical Safety Tests
Focal One complies IEC 62353 regulation about on field Electrical Safety Tests (E.S.T.)
(reminder: IEC 60601-1 regulation is about in factory EST)
E.S.T. are dedicated to patient, operator and FSE safety.
Performing E.S.T. are strongly recommended even if local regulation does not requires it.
To perform EST, follow instruction of latest version of the instruction
TMS511958 or INS-360.
Report result in checklist:
For mobile unit installation TMS511951 or FMU-225
For preventive or curative maintenance TMS511958 or FMU-224
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 5: Error Management
For all software modules, the triggered alarms are coded as follows (only numbers:
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update
W
.
.
X
YY
ZZ

## Slide 6: Error Management
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update
W
.
.
X
YY
ZZ
Were the number mean:
Level of alarm severity, 3 levels are possible:
  Level 1: This level corresponds to a user information that is displayed without an explicit number, and with a level 1 alarm icon. It indicates an event that is easy to solve. The alarm is still logged in the error tracking file.
  Level 2: Non-blocking alarm: Possible actions are "Retry, Continue or Ignore, Cancel or Restart". It is displayed with the error number and a level 2 alarm icon.
  Level 3: Blocking alarm: No possibility to continue treatment, the only option is abandonment. It is displayed with the error number and a level 3 alarm icon.
W

## Slide 7: Error Management
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update
W
.
.
X
YY
ZZ
Were the number mean:
X
Software progress phase at the time of the alarm:
  0: Login phase
  1: Review phase of the treatment database
  2: Initialization phase of the treatment
  3: Phase selection and editing patient data and treatment
  4: Insertion phase of the probe in the patient
  5: Volume acquisition phase
  6: Definition phase of the area to be treated (treatment limits)
  7: Definition phase of lesions on cross sections
  8: Treatment phase (shooting generation)
  9: Zone End Phase (data backup, and move)
  10: End of treatment phase
  100: special phase that groups all the previous phases, but the error is generated with the use of the software FoneSetup

## Slide 8: Error Management
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update
W
.
.
X
YY
ZZ
Were the number mean:
YY
Alarm group: the group refers to a specific module:
  10: Alarms related to the Amplifier
  11: Cooling system group alarms
  12: General Safety Alarms
  13: Alarms related to the detection of the rectal wall
  14: Alarms related to the probe holder (MOT)
  15: Alarms related to the DF probe
  16: Alarms related to the U/S scanne
  17: Matrox board Alarms
  18: Alarms related to file errors
  19: Alarms related to the management of the Focalpak code
  50: Alarms related to general functions of the MEP_Fone
  51: Reasons for canceling treatment
  52: Alarms related to hard disk management

## Slide 9: Error Management
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update
W
.
.
X
YY
ZZ
Were the number mean:
Number of the alarm in the group.
ZZ

## Slide 10: Error Code : Group 10 Amplifier
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update
Alarm | Error message | 
1001 | Measured power too high (higher than 16W above requested for a requested power above 16W) | Can be solved by probe calibration
1002 | Measured power too low (lower than 16W above requested for a requested power above 16W) | 
1003 | Active power < 90% of the required power | 
1004 | Active power > 110% of the required power | 
1005 | Reflected power higher than 10% the direct power | Check good connection between probe and module
1006 | Power error on too many channels (difference> 20%) | Can be solved be probe calibration
1008 | Measured phase out of tolerance | Can be solved be probe calibration
1009 | The initialization of the communication with the generator failed | Check COM port (TechFile and Peripheral Manager). Check amplifier is ON, check fuse FU11.

## Slide 11: Error Code : Group 10 Amplifier
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update
Alarm | Error message | 
1010 | End of trajectory not received by MEP. The shot has been stopped. | Possible computer timeout
1011 | Cannot communicate with the Amplifier | Check COM port (TechFile and Peripheral Manager). Check amplifier is ON, check fuse FU11
1012 | Error status integrity amplifier Error with amplifier initialization | Required full message for troubleshooting.
1013 | Amplifier internal error | Required full message for troubleshooting.
1014 | Measured power too high (higher than 1,6W above requested for a requested power below 16W) | Can be solved be probe calibration
1015 | Measured power too low (lower than 1,6W above requested for a requested power below 16W) | 

## Slide 12: Error Code : Group 11 Cooling system
Alarm | Error message | 
1101 | Probe temperature is too low | If improbable temperature reading (-1000°C) inspect power transducer for crack
1102 | Probe temperature is too high | Perform PM test related to cooling to diagnose defective Peltier
1103 | Cooling system temperature is too low (< 2°C) | Check cooling parameter
1104 | Cooling system temperature is too high (> 10°C) | Check cooling parameter
1105 | Cannot communicate with the Cooling System | Check E/S
1106 | Integrity error on Cooling System | Perform software update
1107 | Status error on Cooling System | 
1108 | The cooling system cannot communicate with the probe | 
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 13: Error Code : Group 12 General safety
Alarm | Error message | 
1201 | Brake failure on probe holder | Monitor magnetic brake feedback
1202 | The patient has moved | If not applicative alarm, probable damage to patient movement detector.
1250 | No 24V: Check if the emergency stop is released | If not user regaled, inspect RL7 and RL8 (burnt marks on relay or its socket).
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 14: Error Code : Group 13 Rectal Wall detection
Alarm | Error message | 
1301 | Rectal wall not detected | Applicative : modify U/S parameter with the US+ dashboard Inspect imaging transducer for shock.
1302 | Rectal wall position error Difference between the recorded rectal wall distance -measured distance during treatment > max deviation before correction | Applicative only
1303 | Rectal wall position error Difference between the recorded rectal wall distance -measured distance during treatment > max deviation after adjustment | Applicative only
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 15: Error Code : Group 14 Probe holder
Alarm | Error message | 
1401 | Cannot communicate with the Motor System | Check E/S and wiring
1402 | Integrity error in a stepper motor board | Perform software update
1403 | Axis status are not the one required | 
1450 | At least one button is pressed on probe holder | Check probe holder keypad in Maintenance
1451 | Communication initialization error on motor | Check E/S and wiring integrity
1452 | 24V2 check failure on motor board | Check E/S
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 16: Error Code : Group 14 Probe holder X axis
Alarm | Error message | 
1410 | Communication timeout with motors | Check motor and coder wiring. Check micrometric ruler integrity Perform PM inspection including random movements
1411 | Actual position is beyond requested position on | 
1412 | Requested position has not been reached on time | 
1413 | Probe movement is impossible. The motor has reached the mechanical stop | 
1414 | Motor position and controled position are different | 
1415 | Command stop displacement failed | 
1416 | Motor position beyond requested position during displacement | 
1417 | Centring not successfully done | 
1418 | Missed encoder pulses | 
1419 | No move detected | 
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 17: Error Code : Group 14 Probe holder Y axis
Alarm | Error message | 
1420 | Communication timeout with motors | Check motor and coder wiring. Check micrometric ruler integrity Perform PM inspection including random movements
1421 | Actual position is beyond requested position on | 
1422 | Requested position has not been reached on time | 
1423 | Probe movement is impossible. The motor has reached the mechanical stop | Applicative: slopped rectal wall due to patient position
1424 | Motor position and controled position are different | Check motor and coder wiring. Check micrometric ruler integrity Perform PM inspection including random movements
1425 | Command stop displacement failed | 
1426 | Motor position beyond requested position during displacement | 
1427 | Centring not successfully done | 
1428 | Missed encoder pulses | 
1429 | No move detected | 
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 18: Error Code : Group 14 Probe holder Θ axis
Alarm | Error message | 
1430 | Communication timeout with motors | Check motor and coder wiring. Perform PM inspection including random movements
1431 | Actual position is beyond requested position on | 
1432 | Requested position has not been reached on time | 
1433 | Probe movement is impossible. The motor has reached the mechanical stop | 
1434 | Motor position and controled position are different | 
1435 | Command stop displacement failed | 
1436 | Motor position beyond requested position during displacement | 
1437 | Centring not successfully done | 
1438 | Missed encoder pulses | 
1439 | No move detected | 
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 19: Error Code : Group 14 Probe holder Z axis
Alarm | Error message | 
1440 | Communication timeout with motors | Check motor and coder wiring. Check micrometric ruler integrity Perform PM inspection including random movements
1441 | Actual position is beyond requested position on | 
1442 | Requested position has not been reached on time | 
1443 | Probe movement is impossible. The motor has reached the mechanical stop | 
1444 | Motor position and controlled position are different | 
1445 | Command stop displacement failed | 
1446 | Motor position beyond requested position during displacement | 
1447 | Centring not successfully done | 
1448 | Missed encoder pulses | 
1449 | No move detected | 
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 20: Error Code : Group 15 Probe
Alarm | Error message | 
1501 | Integrity error on Probe | Check probe connection, check transducer intergrity
1502 | Cannot communicate with the Probe | Check probe connection, check transducer intergrity
1503 | The probe has not been calibrated on this machine | Mismatch between calibrated device S/N, perform probe calibration.
1508 | The probe power has not been calibrated | Partial probe calibration, restart probe calibration for all focals.
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 21: Error Code : Group 16 U/S scanner
Alarm | Error message | 
1601 | Cannot communicate with Ultrasound Scanner | Check U/S scanner is ON, probe has been connected before initialization. No red screen displayed. Remove U/S cover and try connecting probe to another port.
1602 | Problem while launching Ultrasound Scanner Interface software | Re-install software
1603 | Error CRC on U/S Scanner interface software. | Re-install software
1604 | Scanner can't go into running mode | Check U/S scanner is ON, probe has been connected before initialization. No red screen displayed. Remove U/S cover and try connecting probe to another port.
1605 | Problem while retrieving ultrasound scanner scale | Check U/S scanner is ON, probe has been connected before initialization. No red screen displayed. Remove U/S cover and try connecting probe to another port.
1606 | Flip state on ultrasound scanner cannot be set properly | 
1607 | The ultrasound scanner probe has been disconnected. | 
1608 | The ultrasound scanner is shutting down. | 
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 22: Error Code : Group 17 Matrox board
Alarm | Error message | 
1701 | Problem with Ultrasound signal acquisition initialization | Check U/S scanner is ON, Check connections
1702 | Problem with Ultrasound Scanner signal acquisition | Check U/S scanner is ON, Check connections
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 23: Error Code : Group 18 File and hard drive
Alarm | Error message | 
1801 | Impossible to save data | 
1802 | Impossible to open file | 
1803 | Impossible to save TreatmentData file, changes won't be saved. | 
1804 | Impossible to open TreatmentData file, the base will be empty | 
1805 | File integrity corrupt. | 
1806 | Focalpak file integrity corrupt | Use CRC.exe to correct corruption
1807 | Impossible to save Focalpak file | 
1808 | Problem while opening Technical Data file | 
1809 | Problem while saving Technical File | 
1810 | Problem while saving the backed up Technical File | 
1811 | Impossible to create hospital directory | 
1812 | Impossible to open Focalpak file | 
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 24: Error Code : Group 19 FocalPak data
Alarm | Error message | 
1901 | Incorrect Focalpak code | Check Focal S/N on box,check typo
1902 | Focalpak code already used | 
1903 | Duration of Focalpak is out of delay | 
1904 | Access denied! Code invalid for this hospital | 
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 25: Error Code : Group 50 MEP
Alarm | Error message | 
5001 | Cannot communicate with the MEP | Check COM port setting in peripheral manager
5002 | Integrity error on MEP (CRC program, CRC EEPROM, Watchdog, Voltage) | CRC program: perform MEP firmware installation. CRC EEPROM: perform Init EEPROM in Tech File Watchdog or Voltage: replace MEP
5003 | 24V is not present on MEP. Please check Emergency Stop is not pressed | If not user regaled, inspect RL7 and RL8 (burnt marks on relay or its socket).
5004 | Impossible to open COM Port for MEP | Check COM port setting in peripheral manager
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 26: Error Code : Group 51 Treatment cancellation
Alarm | Error message | 
5101 | Treatment cancelled at treatment data edition phase | Information only
5102 | Treatment cancelled at acquisition phase | 
5103 | Treatment cancelled at treatment limits definition phase | 
5104 | Treatment cancelled at transversal localization phase | 
5105 | Treatment cancelled during firing phase | 
5106 | Treatment cancelled because initialization failed | 
5107 | Abnormal treatment termination | 
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 27: Error Code : Group 52 Hard Drive
Alarm | Error message | 
5201 | WARNING : The D:\ disk is running out of space | 10GB remaining warning: delete useless patients (EDAP hospital)
5202 | Not enough available space on the disk to run a treatment. | 5 GB remaining warning: delete useless patients (EDAP hospital)
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update

## Slide 28: Software update
Preventive Maintenance
Error code
Electrical Safety Tests
Error Management
Software Update
Follow instruction of service manual, Service Bulletin, DSR or DII to perform:
MEP firmware update
Motor board firmware update
Cooling board firmware update
Amplifier (mind ESD precaution when handling master board)
Computer software update
Country kit installation